## Slide 1: Report on Electrical Characterisation of 1-layer MoS2 thin film transistors under X-ray radiation (17/11-22/11)
The transistors were fabricated by Adrian Tamayo at the University of Strasbourg. The electrical characterization was performed by Mikhail Bandurist in collaboration with Ilaria Fratelli and professor Laura Basiricò at the University of Bologna.

## Slide 2: Samples structure and characterization
MEASUREMENTS
n-type semiconductor
Low stability in aire, storage in glove box
Electrical measurements were performed reaching VDS = 1 V, and VGS = -60 to 60 V
Delay 100 ms and ΔVGS = 0.5 V
Top contact / Bottom gate
Note: If you need, you can scratch the surface of the substrates to access the gate.
Au/Cr 45/3 nm
MoS2
BCB layer 50 nm
SiOx 278 nm
Si

## Slide 3: Table (Samples 1-8)
Sample_device | W (um) | L (um) | MoS2 thick (nm) | ON/OFF (#/10^6) | SS (V/dec) | Vth (V) | mu (cm^2/(V*s)) | Sensitivity per Volume (C/(Gy*m^3))
S1_d1 | 19.8 | 2.6 | 4.5±0.5 | 0.22 | 4.32±0.16 | -8.75±0.11 | 0.2156±0.0013 | No drain current
S1_d2 | 11.6 | 2.3 | 5.4±0.7 | 0.04 | 12.67±0.08 | 6.87±0.13 | 0.01684±0.00010 | 
S1_d3 | 21.2 | 2.7 | 6.0±0.6 | 3.23 | 2.71±0.05 | -32.1±-.7 | 0.0049±0.0002 | No peaks, steps downwards
S1_d4 | 17.6 | 2.3 | 2.6±0.3 | 0.01 | 8.70±0.12 | 4.6±0.3 | 0.01091±0.00002 | 
S1_d5 | 14.1 | 2.7 | 2.9±0.7 | No current | No current | No current | No current | 
S4_d1 | 13.6 | 2.3 | 2.4±0.3 | 11.30 | 2.85±0.04 | -43.0±0.6 | 0.478±0.012 | No drain current
S5_d1 | 21.8 | 2.3 | 8.9±0.9 | No current | No current | No current | No current | 
S5_d2 | 20.9 | 2.2 | 5.8±0.9 | 0.04 | 5.085±0.26 | -4.62±0.08 | 0.01113±0.00006 | 
S5_d3 | 16.4 | 2 | 2.8±0.4 | No current | No current | No current | No current | 
S5_d4 | No data | No data | 3.1±0.5 | No current | No current | No current | No current | 
S6_d1 | 24.5 | 2.3 | 8.3±0.9 | 14.20 | 4.98±00.10 | -15.61±0.07 | 0.96±0.007 | No peaks, steps downwards
S6_d3 | 21.1 | 2.1 | 10.8±0.6 | No current | No current | No current | No current | 
S6_d5 | 14.7 | 2.1 | 6.3±0.6 | 20.70 | 4.66±0.10 | -6.30±0.10 | 0.26280±0.0014 | No peaks are seen
S7_d1 | 10 | 1.9 | 2.8±0.4 | 3.18 | 5.92±0.07 | -3.2±0.2 | 0.0483±0.0005 | No peaks are seen (only small steps downwards)
S7_d4 | 13.8 | 2.2 | 3.8±0.4 | 0.00 | 6.5±0.3 | -9.23±0.08 | 0.00919±0.00004 | 
S7_d5 | 32.1 | 1.3 | 5.7±0.3 | No current | No current | No current | No current | 
S8_d1 | 8.2 | 2 | 6.2±0.6 | 0.02 | 9.57±0.12 | 10.76±0.16 | 0.2443±0.0016 | No peaks are seen
Common features: • Batch 2 • No encapsulation • Thickness of SiO2 insulator layer = 278 nm
• Capacitance per unit area Ci = 0.0124 (uF/cm^2)

## Slide 4: Table (Samples 9-14)
Sample_device | W (um) | L (um) | MoS2 thick (nm) | ON/OFF (#/10^6) | SS (V/dec) | Vth (V) | mu (cm^2/(V*s)) | Sensitivity per Volume (uC/(Gy*cm^3))
S9_d2 | 20.6 | 2.4 | 7.8±0.7 | Not reliable | 4.03±0.03 | Not reliable | 0.028581272 | 
S9_d3 | 39.6 | 2.2 | 10.8±1.4 | 0.82 | 4.19±0.05 | -14.14±0.09 | 0.784±0.006 | Short circuit
S10_d1 | 12.4 | 3 | 3.9±0.5 | No contact between the sample and the tips | No contact between the sample and the tips | No contact between the sample and the tips | No contact between the sample and the tips | 
S10_d2 | 16.6 | 2.6 | 2.6±0.3 | 0.50 | 10.2±1.0 | 22.3±0.8 | 0.0197±0.0006 | 
S10_d3 | 27.8 | 3.3 | 8.7±0.4 | 0.21 | 2.54±0.05 | -28.37±0.08 | 0.713±0.004 | No drain current
S10_d5 | 14.4 | 3.2 | 8.0±0.5 | No current | No current | No current | No current | 
S10_d6 | 14.8 | 4.4 | 9.4±0.5 | 0.01 | 8.50±0.17 | -45.0±0.4 | 0.329±0.005 | Short circuit
S11_d1 | 14.8 | 3.3 | 3.6±0.4 | 0.00 | 13.9±0.10 | 20.2±0.7 | 0.0309±0.0008 | 
S11_d3 | 8.6 | 2.9 | 9.2±0.8 | 6.44 | 3.66±0.07 | -25.02±0.9 | 0.3481±0.0017 | No sense
S12_d1 | 9.9 | 2.9 | 4.8±0.6 | No current | No current | No current | No current | 
S12_d2 | 13.2 | 3.1 | 2.4±0.3 | 0.07 | 6.50±0.09 | 5.1±0.3 | 0.1231±0.0019 | Too high VGate for reaching photoconductive gain (VGate=50V)
S12_d3 | 20 | 3.8 | 3.3±03 | 0.00 | 8.62±0.11 | -8.11±0.09 | 0.0677±0.0004 | (3.8±0.2)*10^11(VGate=30V) (3.1±1.0)*10^11 (VGate=40V) (-30±7)*10^11 (VGate=50V) – NEGATIVE SENSITIVITY!
S12_d4 | 18.5 | 3.1 | 4.8±0.3 | 1.08 | 6.411±0.004 | -10.20±0.10 | 0.0331±0.0003 | (-1164±73)*10^11 (VGate=50V) NEGATIVE SENSITIVITY!
S14_d1 | 20.1 | 3.5 | 4.8±0.6 | 9.18 | 4.18±0.11 | -10.31±0.10 | 0.353±0.003 | No signal is seen
S14_d2 | 15.5 | 3.2 | 2.4±0.3 | 15.80 | 2.83±0.03 | -42.1±0.5 | 0.288±0.006 | No signal is seen
S14_d3 | 21.1 | 3.7 | 3.3±0.3 | 12.10 | 3.58±0.03 | -13.22±0.05 | 0.1267±0.0003 | (6.5±0.7)*10^11 (VGate=50V)
S14_d4 | 16.4 | 3.2 | 4.8±0.3 | 15.70 | 2.67±0.03 | -55.3±0.7 | 0.274±0.007 | (-123±24)*10^11 (VGate=1V) NEGATIVE SENSITIVITY!
Common features: • Batch 2 • No encapsulation • Thickness of SiO2 insulator layer = 278 nm
• Capacitance per unit area Ci = 0.0124 (uF/cm^2)

## Slide 5: X-ray measurements conditions
X-ray tube with Be window.
Distance between the samples and the Be window ≈30 cm.
Height (both the Be window and the samples) ≈ 22.5 cm.
The samples were placed in the Nextron box with tips.
During all the measurements the Nextron box with the samples was covered by Al foil.
The gate voltage was applied in the range between -50V and +50V.

## Slide 6: Common tendency in most samples
Long (up to 10 minutes) relaxation of the drain current, that could be caused by high concentration of traps in the active channel.
When the X-ray radiation is ON, IDrain starts to decrease more faster than when the radiation is OFF. The result is similar to steps that are directed downwards.
Only two samples have devices that perform positive peaks under X-rays: Sample12 (S12_d3) and Sample14 (S14_d1, S14_d3 and S14_d4).
Some samples have working transfer characteristics, however, they show no signal of any kind even at ITube=500uA.

## Slide 7: Experimental setup
(a)
(b)
(c)
Fig: (a) Experimental setup; (b) X-ray tube control unit; (c) SMU;
(d) Nextron box.
(d)

## Slide 8: Calibration Plot (distance 32 cm from the chassis)

## Slide 9: Sample1_device1
Device 1
Device 2
Device 3
Device 4
Device 5

## Slide 10: Sample1_device1
IDrain shows no signal, probably, bad connection with Nextron.

## Slide 11: Sample1_device3

## Slide 12: Sample1_device3
The initial trfSat is good, the final one is also fine, however, the IGate shows significant leakage.

## Slide 13: Sample1_device3
Long relaxation (high concentration of traps?).
When the X-ray radiation is ON, IDrain starts to sharply decrease, when X-rays are OFF, IDrain decreases more slowly.
At ITube=300uA the noise is too high.

## Slide 14: Sample4_device1
Device 1
Device 2
Device 3
Device 4

## Slide 15: Sample4_device1
IDrain is almost zero, it seemed that the Au electrodes were too damaged.

## Slide 16: Sample6_device1
Device 2
Device 3
Device 4
Device 5

## Slide 17: Sample6_device1
Initial and final trfSat are almost the same and have small hysteresis.

## Slide 18: Sample6_device1
At high ITube, when the X-ray radiation is ON, IDrain starts to sharply decrease. When X-rays are OFF, IDrain decreases more slowly.
At ITube=200uA the tendency is lost.

## Slide 19: Sample6_device5

## Slide 20: Sample6_device5
The initial trfSat is relatively fine, significant hysteresis is present.
IDrain ∝ 10-6 A at VGate=20V.
At ITube=500uA the signal is already unseen, so I stopped the measurement and changed the sample.

## Slide 21: Sample7_device1

## Slide 22: Sample7_device1
At initial trfSat significant hysteresis is present.
IDrain ∝ 10-6 A at VGate=35V.
At ITube500uA the signal was already unseen, so I stopped the measurement and changed the sample. At first peak it is seen that IDrain decreases more sharply when X-ray radiation is ON.

## Slide 23: Sample8_device1

## Slide 24: Sample8_device1
The initial trfSat is realtively fine, significant hysteresis is present.
IDrain ∝ 10-6 A at VGate=22V.
At ITube=500uA the signal was already unseen, so I stopped the measurement and changed the sample.

## Slide 25: Sample9_device3

## Slide 26: Sample9_device3
SHORT CIRCUIT BETWEEN GATE AND DRAIN/SOURCE. Probably it was created by the silver paste intercalating through the insulator.

## Slide 27: Sample10_device3

## Slide 28: Sample10_device3
IDrain is almost zero, the trfSat was also verified in horizontal position. Probably, the Au electrodes were too damaged.

## Slide 29: Sample10_device6

## Slide 30: Sample10_device6
SHORT CIRCUIT BETWEEN THE DRAIN AND THE GATE. Probably it was caused by the silver paste intercalating in the insulator.

## Slide 31: Sample11_device3

## Slide 32: Sample11_device3
NO SENSE.

## Slide 33: Sample12_device2

## Slide 34: Sample12_device2
Significant hysteresis is present.
Drain current reaches microAmps only at 50V, which is too high voltage for reaching photoconductive gain.

## Slide 35: Sample12_device3

## Slide 36: Sample12_device3
At initial trfSat significant hysteresis is present.
IDrain ∝ 10-6 A at VGate=25V.
When the X-ray radiation is ON, POSITIVE peaks in IDrain appear (unlike in previous samples).
To calculate the height of the peaks, each set of peaks was fitted with a normalizing function (polynomial of order 2, exponential or linear).
For each set of peaks the average photocurrent was calculated. After that the dependence of the average photocurrent on the dose rate (related to ITuber) was fitted with a linear function. The slope of the fit determined the sensitivity of the device.
Comments:

## Slide 37: Sample12_device3
At VGate=30V the sensitivity is higher than at VGate=40V.
At VGate=50V the peaks start to decrease with increasing Dose rate (by increasing ITube). As a result, the sensitivity appears to be negative.
Sensitivity/Volume (uC/(Gy*cm3)):
VGate=30V -> (3.8±0.2)*1011
VGate=40V -> (3.1±1.0)*1011
VGate=50V -> (-30±7)*1011 NEGATIVE!

## Slide 38: Sample12_device4

## Slide 39: Sample12_device4
Initial trfSat is relatively fine, hysteresis is present.
IDrain ∝ 10-6 A at VGate=35 V.
Here the peaks also appear but they are NEGATIVE and bigger than in Sample12_device3.

## Slide 40: Sample14_device1

## Slide 41: Sample14_device1
Initial trfSat is relatively fine, small hysteresis is present.
IDrain ∝ 10-6 A at VGate=20V.
Some tendency is seen only at I_Tube=500uA, POSITIVE peaks appear.

## Slide 42: Sample14_device2

## Slide 43: Sample14_device2
Initial and final trfSat are quite good with small hysteresis.
IDrain ∝ 10-6 A at VGate=-2V.
At VGate=1V IDrain shows no tendency, when the radiation is ON, IDrain decreases more sharply than in the dark.

## Slide 44: Sample14_device3

## Slide 45: Sample14_device3
Both initial and the final trfSat have significant hysteresis.
IDrain ∝ 10-6 A at VGate=42V.
Peaks are present and POSITIVE but very small (∝10-9A) . Almost no noise.
At 40 V peaks were quite small (∝10-9A) already at ITube=500uA, so I stopped here and increased VGate to 50V.

## Slide 46: Sample14_device3
The magnitude of the sensitivity per unit volume is similar to Sample12_device3 at VGate=30V and 40V (at 50V the value is negative): ∝ 1011 uC/(Gy*cm3).
Sensitivity/Volume (uC/(Gy*cm3)):
VGate=50V -> (6.5±0.7)*1011

## Slide 47: Sample14_device4 (15/11)

## Slide 48: Sample14_device4 (15/11)
Both initial and final trfSat are fine with almost no hysteresis.

## Slide 49: Sample14_device4 (15/11)
At VGate=1V the peaks are POSITIVE, however, they diminish as ITube increases.
At VGate=2V large negative peaks appear. The same tendency happens at VGate=3V and 5V.

## Slide 50: Sample14_device4 (22/11)

## Slide 51: Sample14_device4 (22/11)
Initial trfSat is quite good with small hysteresis.
IDrain ∝ 10-6 A at VGate=-20V.
At VGate=-20V when the radiation is ON, IDrain starts to sharply decrease with no peaks but STEPS.
At VGate=1V IDrain starts to sharply decrease with almost no peaks or steps present.

## Slide 52: Comparison with the data from literature:
Alberto Taffelli, Max Heyl, Matteo Favaro, Sandra Dirè, Lucio Pancheri, Emil J. W. List-Kratochvil, Alberto Quaranta, Giovanni Ligorio; Demonstrating the high sensitivity of MoS2 monolayers in direct x-ray detectors. APL Mater. 1 August 2023; 11 (8): 081101. https://doi.org/10.1063/5.0151794
Fig: A schematic cross-section of the final device exposed to x ray and connected to an external source meter unit.
Main parameters:
Sensitivity/volume = 108 – 109 uC/(Gy*cm3)
Sensitivity/area = 10 – 100 uC/(Gy*cm2)
Deposition method – exfoliation method
Polysterene (PS) film th ∝ 1 um
Au electrodes th = 100 nm
Active channel area = 0.03 mm2
Active channel th = 0.9 nm
Results from Batch2 (S12_d3 and S14_d3):
S12_d3 Sensitivity/volume = (3.8±0.2)*10^11 uC/(Gy*cm3) (VGate=30V)
S12_d3 Sensitivity/volume = (3.1±1.0)*10^11 uC/(Gy*cm3) (VGate=40V)
S14_d3 Sensitivity/volume = (6.5±0.7)*10^11 uC/(Gy*cm3) (VGate=50V)